INFUSION PUMP SAFETY SIMULATOR

Risk Framework & Scoring Specification

Version 1.0  ·  QR Seed Pilot Study

University Research  ·  Confidential

This document specifies the 4-layer composite risk scoring framework used by the Infusion Pump Safety Simulator to generate AI training labels for medical device safety and security research.

1. Overview

The risk framework assigns a composite risk score (0.0–1.0) and letter grade (A+ through F) to each simulated nurse–pump interaction session. The score drives the risk_label field (low / medium / high) in every TrainingRecord, which serves as the ground-truth label for AI model training.

The framework is deliberately layered — four independent sub-scores assess different concerns, then combine with fixed weights into a single composite. This decomposition allows the AI model to be audited per layer rather than as a black box.

1.1  Research Context

Funding: QR Seed Pilot Study — "An AI-Empowered Safety and Security Ranking System for Infusion Pump Medical Devices."

Primary reference: Cauchi et al. (2011) "Towards Dependable Number Entry for Medical Devices", EICS4Med Workshop. CHI-MED Project.

Clinical basis: Thimbleby & Cairns (2010) J. Royal Society Interface 7(51):1429–1439. Formal model: PVSio-web demos/AlarisGP.

1.2  Devices Covered

| Device Name | model_id | Class | Notes |
| --- | --- | --- | --- |
| BD Alaris GP | alaris_gp | Large-volume volumetric pump | Primary simulator target |
| B. Braun Infusomat Space | braun_infusomat | Large-volume infusion pump | SpaceCom2 / DERS |
| Graseby 3100 | graseby_3100 | Syringe driver (legacy) | Baseline — no safety systems |
| Baxter Sigma Spectrum | sigma_spectrum | Large-volume infusion pump | Phase 2 stub |
| ICU Medical Plum 360 | plum_360 | Large-volume infusion pump | Phase 2 stub |
| Fresenius Kabi Agilia VP | agilia_vp | Large-volume infusion pump | Phase 3 stub |
| CADD-Solis PCA | cadd_solis | PCA / ambulatory pump | Phase 3 stub |

2. Four-Layer Risk Architecture

The composite score is computed from four independent layers. Each layer answers a distinct clinical or technical question. Separating them allows procurement teams, clinical governance, and IT security to act on the layer most relevant to their remit.

| Layer 0 Design Weight: 20% | How safely was this device engineered? Fixed per device model — CVE count, firmware signing, guardrail tiers, VTBI support, anti-freeflow, FDA recalls, display quality. Does NOT change between sessions. Procurement / engineering concern. Source file: deviceDesign.ts |
| --- | --- |

| Layer 1 Interaction Weight: 30% | How well did the nurse program the device? Captures dose error magnitude, direction reversals (corrections), boundary hits, and slow entry time. Source: Cauchi et al. (2011) interaction error taxonomy. Source file: labellingRules.ts (R01, R02, R08, R10, R14, R15, R21) |
| --- | --- |

| Layer 2 Configuration Weight: 25% | Was the session set up safely? Captures guardrail blocks/overrides, MANUAL mode use, VTBI absence, and KCl above limits. Source: ISMP High-Alert Medications; IEC 60601-2-24:2012. Source file: labellingRules.ts (R03, R04, R07, R11, R12, R13) |
| --- | --- |

| Layer 3 System Weight: 25% | What was the device's physical and cyber state? Captures firmware CVE, overdue maintenance, low battery, disconnected gateway, high config drift, and recent occlusion alarms. Source: FDA Infusion Pump Safety guidance; CISA ICS-CERT advisories. Source file: labellingRules.ts (R05, R06, R16–R20) |
| --- | --- |

3. Composite Score Calculation

3.1  Composite Formula

The composite score combines the four layer sub-scores using fixed weights:

composite = 0.20 × design + 0.30 × interaction + 0.25 × configuration + 0.25 × system

Equation 1 — Composite score formula (values in range 0.0–1.0, clamped to 1.0)

The composite_score is the primary feature used for final risk grading. All four layer sub-scores are themselves clamped to [0.0, 1.0] before being combined. The composite is also clamped to [0.0, 1.0] after summation.

3.2  Layer Sub-score Algorithm

Each layer sub-score is computed from the set of risk rules that fired within that layer using the following algorithm (the same algorithm is used for all layers):

layer_score = max(fired_rule_scores) + 0.1 × Σ(remaining_rule_scores)

Equation 2 — Layer sub-score algorithm (clamped to 1.0)

Rationale for this algorithm:

The dominant rule (highest score) contributes its full value.

Each additional rule contributes only 10% of its score, modelling diminishing returns.

This prevents trivial combinations of low-severity rules from producing the same score as a single critical rule.

The cap at 1.0 ensures the score always remains a probability-like value.

3.3  Design Score (Layer 0) — Penalty-based Method

Layer 0 uses a different sub-score algorithm — a direct penalty sum rather than a rule-fired approach — because the inputs are continuous properties of the device profile (CVE count, CVSS score, recall count, etc.) rather than binary session events.

design_score = Σ(applicable_penalty_weights) [clamped to 1.0]

Equation 3 — Design score penalty sum

Penalty weights (source: deviceDesign.ts — DESIGN_PENALTY_WEIGHTS):

| Factor | Penalty | Source |
| --- | --- | --- |
| No drug library (DERS absent) | +0.30 | FDA TPLC Guidance Dec 2014 — drug library is the primary DERS mechanism |
| Guardrail tiers = 0 | +0.25 | IEC 60601-2-24:2012 — essential performance for infusion systems |
| Guardrail tiers = 1 | +0.10 | IEC 60601-2-24:2012 |
| Guardrail tiers = 2 | +0.05 | IEC 60601-2-24:2012 |
| Guardrail tiers = 3 | +0.00 | Full advisory + soft + hard = safest tier |
| No VTBI support | +0.15 | IEC 60601-2-24:2012 Section 201.12 — VTBI as essential safety feature |
| No anti-freeflow protection | +0.10 | Generic Infusion Pump Hazard Analysis v1.0 — Hazard H-12 |
| Per CVE (× count, cap 0.40) | +0.04 | CISA ICS-CERT medical advisories |
| Critical CVE (CVSS ≥ 9.0) | +0.15 | NVD CVSS v3.1 severity bands — nvd.nist.gov |
| Firmware unsigned | +0.15 | CVE-2021-33885 (B. Braun CVSS 9.0) — unauthenticated remote firmware replacement |
| Cleartext transmission | +0.08 | CVE-2021-33883, CVE-2022-26390 (Baxter Sigma Spectrum) |
| Per FDA Class I recall (cap 0.24) | +0.06 | FDA MedWatch / MAUDE database — confirmed patient harm |
| Per FDA Class II recall (cap 0.08) | +0.02 | FDA MedWatch — potential harm |
| Bolus max > 5 ml | +0.08 | Amplifies dose error impact |
| Display quality = basic_numeric | +0.08 | FDA Human Factors Engineering Guidance (2011) |
| Display quality = text_lcd | +0.04 | FDA Human Factors Engineering Guidance (2011) |
| Display quality = graphic_lcd | +0.01 | FDA Human Factors Engineering Guidance (2011) |
| Display quality = color_lcd | +0.00 | Best — no penalty |

4. Grade Thresholds and Risk Labels

The composite_score is mapped to an energy-label-style grade (A+ through F) and then to a risk_label (low / medium / high). The grade is the primary human-readable output; the risk_label is the target variable for AI classification.

| Grade | Score Range | risk_label | Interpretation |  |
| --- | --- | --- | --- | --- |
| A+ | 0.00 – 0.10 | LOW | Exceptional — all safety features present, no errors, no CVEs |  |
| A | 0.11 – 0.20 | LOW | Excellent — minor system issues, perfect interaction |  |
| B | 0.21 – 0.35 | LOW | Good — some system or design concerns, clean interaction |  |
| C | 0.36 – 0.50 | MEDIUM | Adequate — notable risks in one or more layers |  |
| D | 0.51 – 0.65 | MEDIUM | Poor — multiple risk factors across layers |  |
| E | 0.66 – 0.80 | HIGH | Very poor — significant design risk and/or session issues |  |
| F | 0.81 – 1.00 | HIGH | Critical — severe design risk, major session error, or both |  |

Note: A single HIGH risk rule firing does not automatically set risk_label = "high" in the composite model. The grade is always derived from the composite_score. However, B. Braun-specific device rules (BB-R02, BB-R05) do force a HIGH label override when they fire, because they represent immediate, session-active threats (network + unsigned firmware, or FDA-recalled device).

5. Risk Rules — Complete Specification

Rules R01–R21 are generic (all devices). B. Braun-specific rules BB-R01–BB-R05 supplement the generic rules for the Infusomat Space model.

5.1  HIGH Risk Rules (R01–R08)

Any single HIGH rule firing is sufficient to set risk_label = "high" via the composite grade.

| ID | Layer | Score | Test Condition | Label | Source |
| --- | --- | --- | --- | --- | --- |
| R01 | L1 Interaction | 1.00 | relative_error ≥ 0.90 | 10× dose error (≥90% relative error) | Thimbleby & Cairns (2010) J. Royal Society Interface 7(51) — 10× dose errors are the most severe category of programming error identified in the dose error taxonomy. |
| R02 | L1 Interaction | 0.85 | relative_error ≥ 0.50 | 50% dose error (≥50% relative error) | Thimbleby & Cairns (2010) — 50% error threshold; ISMP Medication Error Reports. |
| R03 | L2 Configuration | 0.80 | guardrail_blocked = 1 | Hard guardrail limit triggered (blocked from infusing) | IEC 60601-2-24:2012 Section 201.12 — hard limits define the absolute safe operating range; triggering them indicates programming outside the verified therapeutic window. |
| R04 | L2 Configuration | 0.75 | drug_library_used = 0 AND final_rate_ml_h > 200 | MANUAL mode with high rate (>200 ml/h) — no library protection | FDA TPLC Guidance Dec 2014 — MANUAL mode bypasses all DERS; high rates amplify the consequence of any concurrent programming error. |
| R05 | L3 System | 0.70 | firmware_version_risk = 1 | Firmware version has known CVE vulnerability | CISA ICS-CERT medical device advisories. CVE list includes: 6.0.2, 7.1.0, 8.05, 8.1.3 (BD Alaris). CVE-2022-22772 (CVSS 9.8) — unauthenticated remote access. |
| R06 | L3 System | 0.65 | days_since_maintenance > 365 | No maintenance in >1 year | FDA Infusion Pump Safety guidance — maintenance overdue by >1 year is associated with increased hardware failure rate and undetected calibration drift. |
| R07 | L2 Configuration | 0.90 | drug_id = "kcl" AND rate_within_soft_limits = 0 | KCl rate above soft limits — cardiac arrest risk | ISMP High-Alert Medications list — KCl is classified as a high-alert drug. Rapid KCl infusion above soft limits is associated with fatal cardiac arrhythmia. National Patient Safety Agency (UK) Safety Alert 2002. |
| R08 | L1 Interaction | 0.80 | confirmed_incorrect = 1 AND relative_error > 0.10 | Confirmed wrong value — infusion started with >10% error | Cauchi et al. (2011) EICS4Med — confirmed commitment to an incorrect value despite the opportunity to review; models the "confirmation bias" failure mode. |

5.2  MEDIUM Risk Rules (R10–R21)

Any single MEDIUM rule firing (without a HIGH rule) is sufficient for risk_label = "medium". Multiple MEDIUM rules accumulate via the layer sub-score algorithm.

| ID | Layer | Score | Test Condition | Label | Source |
| --- | --- | --- | --- | --- | --- |
| R10 | L1 | 0.40 | relative_error ≥ 0.10 | 10% dose error (≥10% relative error) | Thimbleby & Cairns (2010) — 10% is the clinical significance threshold below which most drug effects are within normal pharmacokinetic variability. |
| R11 | L2 | 0.45 | guardrail_override = 1 | Soft guardrail warning overridden by operator | Cauchi et al. (2011); AAMI TIR12:2010 — guardrail overrides are the primary failure mode identified in DERS implementation studies. |
| R12 | L2 | 0.30 | drug_library_used = 0 | MANUAL mode — no drug library protection | FDA TPLC Guidance — MANUAL mode removes all DERS guardrail protection. |
| R13 | L2 | 0.25 | vtbi_set = 0 | VTBI not set — no infusion completion detection | IEC 60601-2-24:2012 Section 201.12; BD Alaris GP DFU BDDF00535 Issue 4. |
| R14 | L1 | 0.30 | correction_count > 5 | Excessive corrections during entry (>5 direction reversals) | Cauchi et al. (2011) — correction count is the primary behavioural indicator of number entry difficulty correlated with error probability. |
| R15 | L1 | 0.20 | boundary_hit_count > 0 | Boundary hit during rate entry — operator tested device limits | Cauchi et al. (2011) — boundary hits indicate navigation to the extreme of the allowed range, associated with high-error sessions. |
| R16 | L3 | 0.25 | drug_library_age_days > 90 | Drug library outdated (>90 days since update) | BD Alaris GP DFU recommends library updates per hospital formulary change cycle; ISMP recommends quarterly review minimum. |
| R17 | L3 | 0.35 | battery_level_pct < 20 | Battery critically low (<20%) | BD Alaris GP DFU Section 5 — battery warning at <30 min remaining; <20% approximates this threshold. |
| R18 | L3 | 0.30 | network_connected = 0 | Not connected to Gateway — no remote monitoring | FDA Infusion Pump Safety guidance — network connectivity enables dose error reporting, library push, and remote alarm monitoring. |
| R19 | L3 | 0.30 | config_drift_score > 0.50 | High configuration drift (>0.5) from hospital standard | AAMI TIR12:2010 — configuration drift from hospital-standard settings is identified as a systemic risk factor for pumps in ward use. |
| R20 | L3 | 0.40 | recent_occlusion_alarms ≥ 3 | Recent occlusion alarms (≥3 in last 24h) | FDA MAUDE database — repeat occlusion alarms in the same 24-hour period indicate mechanical issues (kinked line, failing pump mechanism). |
| R21 | L1 | 0.20 | entry_time_ms > 60000 | Slow rate entry (>60 seconds) | Cauchi et al. (2011) — entry time >60s correlated with confusion and higher error probability in the number entry study. |

5.3  B. Braun-Specific Rules (BB-R01–BB-R05)

These rules supplement R01–R21 specifically for the B. Braun Infusomat Space. They are implemented in braunRules.ts and applied in addition to the generic rules.

| ID | Severity | Layer | Score | Test | Label | Source |
| --- | --- | --- | --- | --- | --- | --- |
| BB-R01 | HIGH → MEDIUM* | L3 System | 0.75 | firmware_signed = 0 | Firmware not cryptographically signed — systemic background risk | CISA ICSMA-21-294-01 — CVE-2021-33885 (CVSS 9.0). *Reclassified MEDIUM because this is a permanent device property already penalised in Layer 0 design_score. Exploitation requires active network access; without SpaceCom2 connected there is no immediate session-level threat. |
| BB-R02 | HIGH | L3 System | 0.85 | spacecom2_connected = 1 AND firmware_signed = 0 | Networked device with unsigned firmware — actively exploitable this session | CISA ICSMA-21-294-01 — CVE-2021-33882 (CVSS 7.7): missing authentication for network commands. Combined with unsigned firmware = attacker on same network can modify pump parameters without credentials. Immediate session-active threat. |
| BB-R03 | MEDIUM | L2 Configuration | 0.35 | guardrail_advisory_acknowledged = 1 AND guardrail_warning_shown = 1 | Advisory dismissed, then soft limit subsequently triggered | B. Braun IFU — Advisory tier is the earliest warning before soft stop. Dismissing it then crossing the soft limit indicates escalating boundary-pushing behaviour correlated with dose error risk. |
| BB-R04 | MEDIUM | L2 Configuration | 0.30 | bolus_delivered = 1 AND bolus_volume_ml > 1.5 | Bolus volume > 75% of device maximum (>1.5 ml of 2.0 ml max) | B. Braun IFU — Bolus Volume Max 2 ml. For concentrated drugs (KCl 1 mmol/ml, heparin 1000 U/ml), a near-maximum bolus delivers a clinically significant unintended dose. 75% threshold follows standard near-maximum operation margin. |
| BB-R05 | HIGH | L0 Design | 0.90 | recall_batch_affected = 1 | Device in FDA 2023 Class I recall scope — faulty occlusion alarm | FDA Class I Recall Z-0601-2024 (November 2023): faulty occlusion detection software. ~10,600 devices affected. Pump fails to alarm on downstream occlusion at certain flow rates. One confirmed patient death — vasopressor (noradrenaline) interruption without staff alert. |

6. Sources and References

Cauchi et al. (2011). Cauchi M, Curzon P, Blandford A. "Towards Dependable Number Entry for Medical Devices." EICS4Med Workshop @ CHI-MED Project. Cambridge, 2011. Primary source for R08, R10, R14, R15, R21 — interaction error taxonomy.

Thimbleby & Cairns (2010). Thimbleby H, Cairns P. "Reducing Number Entry Errors: Solving a Widespread, Serious Problem." J. Royal Society Interface 7(51):1429–1439. 2010. Source for R01, R02, R10 — dose error magnitude thresholds.

IEC 60601-2-24:2012. International Electrotechnical Commission. "Medical electrical equipment — Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers." IEC 60601-2-24:2012+AMD1:2020. Source for R03, R13 — VTBI and guardrail essential performance requirements.

FDA TPLC Guidance (2014). U.S. Food and Drug Administration. "Infusion Pumps Total Product Life Cycle: Guidance for Industry and FDA Staff." December 2014. Source for R04, R05, R12 — DERS implementation, MANUAL mode risk, firmware CVE.

FDA Infusion Pump Safety. U.S. FDA. "Infusion Pumps." fda.gov/medical-devices/infusion-pumps. Source for R06, R17, R18 — maintenance, battery, network monitoring.

ISMP High-Alert Medications. Institute for Safe Medication Practices. "ISMP List of High-Alert Medications in Acute Care Settings." 2018. www.ismp.org/tools/highalertmedications.pdf. Source for R07 — KCl cardiac arrest risk; R11, R12 — guardrail override consequences.

BD Alaris GP DFU. BD (Becton, Dickinson & CareFusion). "Alaris GP Volumetric Pump — Directions For Use." BD document 1000DF00152 Issue 1 / BDDF00535 Issue 4. Source for device constants (RATE_MIN/MAX, VTBI_MAX, KVO_RATE, BOLUS_RATE), R13, R16.

CISA ICSMA-21-294-01. Cybersecurity and Infrastructure Security Agency. "B. Braun Infusomat Space Large Volume Pump and SpaceStation." ICS Medical Advisory ICSMA-21-294-01. October 2021. CVE-2021-33885 (CVSS 9.0) — unauthenticated firmware modification. CVE-2021-33882 (CVSS 7.7) — missing auth for network commands. CVE-2021-33883 — cleartext data transmission. Source for BB-R01, BB-R02.

CISA ICSMA-23-194-01. Cybersecurity and Infrastructure Security Agency. "BD Alaris Infusion System." ICS Medical Advisory ICSMA-23-194-01. July 2023. 13 vulnerabilities. CVE-2022-22772 (CVSS 9.8) — unauthenticated remote access. Source for R05.

CISA ICSMA-22-251-01. Cybersecurity and Infrastructure Security Agency. "Baxter Sigma Spectrum Infusion Pumps." ICS Medical Advisory ICSMA-22-251-01. September 2022. CVE-2022-26390 (CVSS 7.5) — cleartext transmission. Source for deviceDesign.ts sigma_spectrum.

CISA ICSMA-16-306-01. Cybersecurity and Infrastructure Security Agency. "Smiths Medical CADD-Solis Infusion Pump." ICS Medical Advisory ICSMA-16-306-01. November 2016. CVE-2016-8375 (CVSS 9.9) — patient-accessible memory with credentials in plaintext. Source for deviceDesign.ts cadd_solis.

FDA Recall Z-0601-2024. U.S. Food and Drug Administration. "B. Braun Medical Inc. Infusomat Space and Infusomat fmS Large Volume Pumps." Class I Recall Z-0601-2024. November 2023. ~10,600 devices affected — faulty occlusion detection software. One confirmed patient death (vasopressor interruption). Source for BB-R05.

NPSA Safety Alert 2010. National Patient Safety Agency (UK). "Safer use of injectable medicines." NPSA/2010/RRR013. 2010. Graseby syringe driver safety alert — opioid overdose deaths. Source for graseby_3100 design profile rationale.

AAMI TIR12:2010. Association for the Advancement of Medical Instrumentation. "Infusion System Risk Management and Safety." AAMI TIR12:2010. Source for R11, R12, R19.

Generic Infusion Pump Hazard Analysis v1.0. FDA-funded report. ResearchGate. "Generic Infusion Pump Hazard Analysis." Hazard H-12 — gravity-driven freeflow bolus on set disconnection. Source for NO_ANTI_FREEFLOW penalty in deviceDesign.ts.

PVSio-web Formal Model. Masci P, et al. "PVSio-web: A Tool for Rapid Prototyping Device Interfaces." github.com/pvsioweb/pvsio-web/tree/master/examples/demos/AlarisGP. Demo: pvsioweb.org/demos/AlarisGP. Source for state machine transitions.

7. Device Design Scores — Layer 0 Leaderboard

Pre-computed design scores for all registered device profiles. Lower score = safer design. Sorted safest first.

| Device | model_id | Score | Grade | Key Penalties |
| --- | --- | --- | --- | --- |
| B. Braun Infusomat Space | braun_infusomat | 0.73 | E | 5 CVEs, CVSS 9.0, unsigned firmware, cleartext, 2×Class I, 1×Class II, graphic LCD |
| BD Alaris GP | alaris_gp | 0.88 | F | 13 CVEs, CVSS 9.8, unsigned firmware, cleartext, 3×Class I, 2×Class II, text LCD |
| Fresenius Kabi Agilia VP | agilia_vp | 0.85 | F | 13 CVEs, CVSS 7.5, unsigned firmware, cleartext, 1×Class I, graphic LCD |
| CADD-Solis PCA | cadd_solis | 0.82 | F | 2 CVEs, CVSS 9.9, no anti-freeflow, no KVO, 3×Class I |
| Baxter Sigma Spectrum | sigma_spectrum | 0.67 | E | 4 CVEs, CVSS 7.5, unsigned firmware, cleartext, 3×Class I, color LCD (best display) |
| ICU Medical Plum 360 | plum_360 | 0.38 | C | 0 CVEs, no cleartext, unsigned firmware, 2×Class I, text LCD |
| Graseby 3100 | graseby_3100 | 0.63 | D | No drug library, no guardrails, no VTBI, no anti-freeflow, no KVO, basic numeric display |

Note: Scores are computed at module load time from DEVICE_PROFILES_RAW in deviceDesign.ts and cached in DEVICE_REGISTRY. They are stable across all sessions for a given device model.